## Slide 1
07/01/2016 Claire Prouve - University of Bristol LHCb UK Liverpool
RICH Mirror Alignment
presenting work by Anatoly, Antonis, Chris, Claire, Paras, Roel, and many more

## Slide 2: Overview
RICH mirror alignment strategy
RICH mirror alignment implementation in Run II
Performance in Run II
RICH maintenance and operation

## Slide 3: RICH Mirror Alignment
RICH 1:
RICH 2:
Misaligned mirrors will affect the PID due to incorrectly predicted Cherenkov angle!
Primary mirrors
Primary mirrors
Secondary mirrors
Secondary mirrors
4 primary mirrors
16 secondary mirrors
54 primary mirrors
40 secondary mirrors
Photon Detectors
Photon Detectors
No more Aerogel!
x
x

## Slide 4: Misalignment
Proj. Track position
Real track position
Misalignment
Cherenkov angle resolution / rad
azimuthal angle/ rad
Misaligned mirror:
Aligned mirror:
Identify misalignment:
Misalignments on detector plane
Detector plane
ΔθC(Φ) = θmeas.- θexp.
Cherenkov angle resolution / rad
azimuthal angle/ rad
ΔθC(Φ) = ρy cos(Φ) + ρz sin(Φ)

## Slide 5: Decoupling misalignments
RICH1: easy!
Given secondary mirror only receives photons from one primary mirror.
 Only align secondary mirrors
Mirror-pair: fit misalignment on detector-plane in y, z  2 parameters
need actual misalignment in y,z, for each mirror
 4 parameters
Primary mirrors
Secondary mirrors

## Slide 6: Decoupling misalignments
Mirror-pair: fit misalignment on detector-plane in y, z  2 parameters
need actual misalignment in y,z, for each mirror
 4 parameters
Andrew Cook
RICH2:
Given secondary mirror can receive photons from several primary mirrors.
System of equations linking all mirrors starting from primary mirror 12.
primary mirrors, secondary mirrors

## Slide 7: Alignment procedure
Unambiguous photons will be reflected by the same mirror-pair no matter where along the track they were emitted.
Events
Fill histograms with (Δθ vs. Φ) of unambiguous photons for each mirror- pair
Calculate mirror misalignments for each individual mirror + Update database
CONDDB database
High energy tracks + reconstruct under pion- hypothesis
Determine misalignment on detector plane
new database

## Slide 8: Alignment procedure
Unambiguous photons will be reflected by the same mirror-pair no matter where along the track they were emitted.
Events
Fill histograms with (Δθ vs. Φ) of unambiguous photons for each mirror- pair
CONDDB database
High energy tracks + reconstruct under pion- hypothesis
Determine misalignment on detector plane
new database
Any mirror moved > 0.1 mrad?
Calculate mirror misalignments for each individual mirror + Update database

## Slide 9: Alignment procedure
Unambiguous photons will be reflected by the same mirror-pair no matter where along the track they were emitted.
Events
Fill histograms with (Δθ vs. Φ) of unambiguous photons for each mirror- pair
High energy tracks + reconstruct under pion- hypothesis
Determine misalignment on detector plane
new(er) database
YES!
new database
Calculate mirror misalignments for each individual mirror + Update database

## Slide 10: Alignment procedure
Unambiguous photons will be reflected by the same mirror-pair no matter where along the track they were emitted.
Events
Fill histograms with (Δθ vs. Φ) of unambiguous photons for each mirror- pair
High energy tracks + reconstruct under pion- hypothesis
Determine misalignment on detector plane
new(er) database
Any mirror moved > 0.1 mrad?
new database
Calculate mirror misalignments for each individual mirror + Update database

## Slide 11: Alignment procedure
Unambiguous photons will be reflected by the same mirror-pair no matter where along the track they were emitted.
Events
Fill histograms with (Δθ vs. Φ) of unambiguous photons for each mirror- pair
High energy tracks + reconstruct under pion- hypothesis
Determine misalignment on detector plane
new(er) database
No!  Alignment converged!!!
new database
Calculate mirror misalignments for each individual mirror + Update database

## Slide 12
RICH mirror alignment implementation in Run II

## Slide 13: Run I and Run II
Run I:
Offline, after data taking
with Ganga for the Brunel reconstruction
Alignment applied at the end-of-year reprocessing
Run II:
Online: after HLT1 and before HLT2
Alignment run for each Fill (monitoring mode)
HLT1 lines for alignment
HLT farm (~1700 nodes)
Steered by the online Finite State Machine (FSM)

## Slide 14: HLT farm for Run II
Beat Jost
HLT farm
HLT farm nodes (Analysers)
  Reconstruct data from HLT1 lines
  Make the histograms
Central node (Iterator)
  Receives histograms from Analysers
  Determines misalignment on detector plane
  Determines misalignment of each individual mirror
  Makes new database + decides weather or not to continue

## Slide 15: Finite State Machine
Set up Brunel reconstruction options.
Get default database and place where analyzers will pick it up, etc.
Reconstruct events and fill histograms.
Mirror misalignments + new database.
In case of continue: place new database where the analyzers will pick it up.
Reconstruct events and fill histograms.
configure
configure
CONFIGURING
CONFIGURING
READY
start
RUNNING
PAUSED
start
RUNNING
stop
READY
READY
PAUSED
pause
RUNNING
start
PAUSED
pause
READY
READY
RUNNING
stop
Alignment converged!
Iterator
Analyzer(s)
Run Control

## Slide 16
Performance in Run II

## Slide 17: Performance – RICH2
RICH2 Cherenkov Resolution Stability : 2015
Run I:
<σΔθ> ~ 0.67 mrad
Run II average: <σΔθ> ~ 0.67 mrad
Early measurements
Chris Jones

## Slide 18: Performance – RICH1
RICH1 Cherenkov Resolution Stability : 2015
Chris Jones
Run I:
<σΔθ> ~ 1.62 mrad
Run II average: <σΔθ> ~ 1.65 mrad
Small but clear differences between magnet up and magnet down!
Magnet down
Magnet up
Chris Jones

## Slide 19: PID Performance
Sneha Malde
Δ LL (K – π) > 0
Δ LL (K – π) > 5
Run I:
Run II:

## Slide 20
Maintenance and operation

## Slide 21: Calibrations
Online calibrations at the end of each run.
Refractive Index Calibration:
Hardware sensors monitor pressure and temperature.
Limited precision and does not account for gas mixture changes.
Simple fit to reconstructed-expected Cherenkov angle yields (n-1) scale factor.
HPD images calibration:
Image fit performed for each HPD and used to provide calibration for anode element.
Sobel Filter

## Slide 22: Maintenance and Operation
Aerogel in RICH1 removed
~50% of HPDs in RICH1 replaced and many in RICH2  quite new detector
Gas leak in RICH1: not found but workaround developed
Gas pressure generally well handled
Some incomplete events due to RICH detectors  bug in UKL0 firmware found and fixed
RICHes performed well during the heavy ion runs
Average pp event:
Average heavy ion event:

## Slide 23: Summary
RICH detectors are performing very well
Alignment will be run online for every fill
  PID can be used in HLT2
  Very fast (<2h wrt. ~2 days)
RICH2 reached precision of RunI already
RICH1 needs a bit more work
More optimisation is being done

## Slide 24
Backup

## Slide 25: Cherenkov angle resolution
Limiting factors to Cherenkov angle resolution:

## Slide 26: Misalignment
Δθ [rad]
Φ [rad]
Misaligned mirrors:
Aligned mirrors:
26
Δθ [rad]
Φ [rad]
Signal:
Gaussian
BKG:
2nd order polynomial

## Slide 27: Magnification coefficients
Magnification coefficients: Translate the tilt on the detector plane into actual mirror tilts
Magnification coefficients are calculated new for each iteration:
Introduce 8 rotations: primary and secondary mirrors rotated around ±y and ±z axis respectively
Rotate about 0.3 mrad ( half the resolution of RICH2)
Reconstruct events for each rotation and evaluate the tilts on the detector plane
Need to reconstruct all events 9 times!
Misalignments on detector plane
Mirror tilts

## Slide 28: New HLT Lines
Track pseudorapidity
Track azimuthal angle / rad
Trigger on tracks that will populate the hardest-to-populate mirror-pairs
usually the very outer mirrors
Other tracks in the events will populate the rest
RICH2
Track pseudorapidity
RICH1
RICH2 line:
p > 40 GeV && χ2 < 2 && 2.65 < η < 2.80
(-2.59 < Φ < -2.49) || ( -0.65 < Φ < -0.55) || (0.45 < Φ < 0.65) || (2.49 < Φ < 2.59)
RICH1 line:
p > 10 GeV && χ2 < 2 && 1.6 < η < 2.04
(-2.65 < Φ < -2.3) || (-0.8 < Φ < -0.5) || (0.5 < Φ < 0.8) || (2.3 < Φ < 2.65)
Need to reconstruct ~10 times less events!

## Slide 29: Example RICH1 improvement
First alignment in RunII
Second alignment in RunII