## Slide 1
22/06/2016 Claire Prouve - University of Bristol LHCb Week
Real-time alignment and calibration: the RICH systems
presenting work by Anatoly, Antonis, Chris, Claire, Jibo, Paras and many more

## Slide 2: Overview
RICH calibrations
RICH mirror alignment
  Procedure
  Improvements during Run II
Results

## Slide 3: RICH calibrations
Refractive index calibration:
Hardware sensors monitor pressure and temperature.
Limited precision and does not account for gas mixture changes.
Simple fit to reconstructed-expected Cherenkov angle yields (n-1) scale factor.
HPD image calibration:
Image fit performed for each HPD and used to provide calibration for the anode element.
Fully automatically for every run since beginning of Run II.
Sobel Filter

## Slide 4: RICH mirror alignment
Proj. Track position
Real track position
Misalignment
Cherenkov angle resolution / rad
azimuthal angle/ rad
Misaligned mirror:
Aligned mirror:
Identify misalignment:
Misalignments on detector plane
Detector plane
ΔθC(Φ) = θmeas.- θexp.
Cherenkov angle resolution / rad
azimuthal angle/ rad
ΔθC(Φ) = Θyp,s cos(Φ) + Θzp,s sin(Φ)

## Slide 5: Mirror alignment procedure
Events
Fill histograms with (Δθ vs. Φ) of unambiguous photons for each mirror com- bination.
Select high momentum tracks + reconstruct under pion- hypothesis.
database
Produce new database and verify if the convergence criteria was met.
Fit misalignment on the detector plane and determine individual mirror misalignments.
Alignment
converged!
Replace database and perform another iteration.
YES
NO

## Slide 6: Development during Run II
 | Reco. | Mode | Improvements | Time per alignment
Run I | Ganga | manually |  | Several days
Run II | HLT farm | manually | Migrated to online Dedicated HLT1 lines | ~4 hours
Run II end of 2015 | HLT farm | manually | Using pre-calculated magnification coefficients Improved method for fitting 2D histograms | ~20 minutes
Run II beginning of 2016 | HLT farm | automatic for every fill | L2 regularization method for finding individual mirror tilts New CondDB | ~10 minutes after mag. flip: ~30 minutes for RICH1
Run II end of June 2016 | HLT farm | automatic for every fill | Different convergence criteria for primary/secondary mirrors in y/z Calculation of difference between final alignment and alignment in CondDB | ~10 minutes for each RICH detector

## Slide 7: Stability of tilts – RICH1
Mag Down
Mag Up
Convergence criteria
0.1 mrad

## Slide 8: Stability of tilts – RICH2
Convergence criteria
0.1 mrad

## Slide 9: Stability of Resolution
magnet down
magnet up
magnet down before new DB
Resolutions from refractive index calibration (without HPD image correction).

## Slide 10: Current monitoring
https://lbgroups.cern.ch/rich/calibview.php
https://lbgroups.cern.ch/rich/alignmentview.php
Tilts in y
Tilts in z
RICH2 primary mirror tilts w.r.t. CondDB
RICH2 secondary mirror tilts w.r.t. CondDB
number of secondary mirror

## Slide 11: Summary
RICH calibrations run online for every run
RICH mirror alignment runs online for every fill
Several improvements in stability and speed
Information from frequently run alignments to understand and make improvements
Monitoring plots for the RICH piquet

## Slide 12
Backup

## Slide 13: Cherenkov angle resolution
Limiting factors to Cherenkov angle resolution:

## Slide 14: Fit to 2D histograms
Δθ [rad]
Φ [rad]
Misaligned mirrors:
Aligned mirrors:
14
Δθ [rad]
Φ [rad]
Signal:
Gaussian
BKG:
2nd order polynomial
Means of Gaussians connected by:
ΔθC(Φ) = Θyp,s cos(Φ) + Θzp,s sin(Φ)

## Slide 15: Magnification coefficients
Magnification coefficients: Translate the tilt on the detector plane into actual mirror tilts
Magnification coefficients are calculated new for each iteration:
Introduce 8 rotations: primary and secondary mirrors rotated around ±y and ±z axis respectively
Rotate about 0.3 mrad ( half the resolution of RICH2)
Reconstruct events for each rotation and evaluate the tilts on the detector plane
Need to reconstruct all events 9 times!
Misalignments on detector plane
Mirror tilts

## Slide 16: Disentangling mirror-pairs
Each pair of primary mirror p and secondary mirror s has 2 equations:
Magnification factors individual mirror tilts
RICH1 primary mirrors
RICH1 secondary mirrors
Challenge: not enough information to fully constrain the solution
sss
RICH1: 8 equations for 10 unknowns aaa
Rotation of primary mirrors followed by according rotation of secondary mirrors yields same results s
 Need additional constraint
ΔθC(Φ) = Θyp,s cos(Φ) + Θzp,s sin(Φ)

## Slide 17: Disentangling – until 2016
RICH1:
fix primary mirrors, align secondary mirrors
Primary mirrors
Secondary mirrors
Andrew Cook
primary mirrors, secondary mirrors
17
RICH2:
Left half: system of equations linking all mirrors starting from primary mirror 12.

## Slide 18: Disentangling – from 2016
L2 regularization (ridge regression): d
rewrite problem as
Minimize: least square method L2 regularization term
Advantages:
more stable than the previous method
Stable w.r.t. small statistical fluctuations (unlike L1 regularization)
 Fewer iterations needed to converge
Matrix of
magnification factors
Vector of results of fit to histograms
Vector of individual mirror tilts

## Slide 19: Resolution
magnet down
magnet up
magnet down before new DB
Without HPD image correction: Degradation of resolution over fill.
 Longer fills in 2016

## Slide 20: New HLT Lines
Track pseudorapidity
Track azimuthal angle / rad
Trigger on tracks that will populate the hardest-to-populate mirror-pairs
usually the very outer mirrors
Other tracks in the events will populate the rest
RICH2
Track pseudorapidity
RICH1
RICH2 line:
p > 40 GeV && χ2 < 2 && 2.65 < η < 2.80
(-2.59 < Φ < -2.49) || ( -0.65 < Φ < -0.55) || (0.45 < Φ < 0.65) || (2.49 < Φ < 2.59)
RICH1 line:
p > 10 GeV && χ2 < 2 && 1.6 < η < 2.04
(-2.65 < Φ < -2.3) || (-0.8 < Φ < -0.5) || (0.5 < Φ < 0.8) || (2.3 < Φ < 2.65)
Need to reconstruct ~10 times less events!